#### ANNEX WP1.A: Minutes concerning WP1.A Solar Collectors

The 5th Solar Keymark meeting in Lisbon, March 5-7, 2003

Minutes taken by Åsa Wahlström SP (AAW).

# Implementation of test procedures and update of time schedule

To label solar collectors with the Solar Keymark test must have been performed according to EN12975-1, paragraph 5.2 Required tests a)-j).

- Six laboratories are ready with the accreditation.

- Four laboratories are planning to get accreditation during 2003.

- All laboratories are ready with the implementation of the tests.

- DTI has implemented the tests but has no planes for accreditation.

- CSTB is planning to be accredited only for test h) Thermal performance

- TNO are considering to get accredited for several selected tests. Additional tests must be performed at another laboratory.

- Five laboratories will exclude test i) Freeze resistance test. This test is, however, only for collectors that are claiming to be freeze-resistant. The laboratories therefore can offer complete tests for Solar Keymark for all other collectors while collectors that are claiming that they are freeze resistant must do test i) at another laboratory.

The implementation and accreditation plans are shown in Table WP1.A.1 and Figure WP1.A.1.

Figure WP1.A.1 	Accreditation of test procedure according to EN12975-1,

Solar Collectors

Table WP1.A.1 	Implementation and accreditation of test procedure.

EN12975-1, 2 Solar Collectors

| Laboratory | Implemented | Accreditation | Tests planning for accreditation stated in EN 12975-1 § 5.2 a) – j) |
| --- | --- | --- | --- |
| Arsenal (Austria) | Ready | Ready | All |
| CSTB (France) | Ready | 2003-12-31 | h) Thermal performance |
| Demokritos (Greece) | Ready | Ready | Excluded test: i) Freeze resistance |
| DTI (Denmark) | Ready | No | Excluded tests: i) Freeze resistance |
| ENEA (Italy) | Ready | 2003-04-30 | Excluded tests: i) Freeze resistance |
| INETI (Portugal) | Ready | 2003-12-31 | Excluded test: i) Freeze resistance |
| INTA (Spain) | Ready | 2003-12-31 | All |
| ITW (Germany) | Ready | Ready | All |
| ITC (Spain) | Ready | Thermal performance accreditation ready. Reliability tests accreditation 2003-03-31 | Excluded tests: i) Freeze resistance |
| IZES (Germany) | Ready | Ready | All |
| SP (Sweden) | Ready | Ready | All |
| SPF (Switzerland) | Ready | Ready | All |
| TNO (The Netherlands) | Ready | In consideration | b), c), d), e), f), g) and j) |

###### Round Robin of uncertainty of test result

Stephan Fischer (ITW) presented a “Round Robin” of uncertainty in collector efficiency parameters. The objectives with the Round Robin are:

- to share the already gained experience

- to agree on a consolidated procedure to calculate the uncertainties in the collector efficiency parameters

The first round of the Round Robin had the following steps:

- Determination of the maximum allowed uncertainty

- Determination of the type of uncertainty (A or B)

- Assumed distribution of the uncertainty

- Agreement of the standard uncertainty used for further calculations.

Six laboratories participated in the first round. The results can be found tabulated in

Annex WP1.A.A, “Uncertainty power point presentation” by Stephan Fischer.

Notable results so far is that:

- The collector outlet temperature te has no stated uncertainty in the standard. Most participants assumed that it should have the same uncertainty as the inlet temperature, tin.

- Class 1 pyranometer should have a standard uncertainty of 1 %, which will give a maximum uncertainty of 2.45 %.

The next step will be to get an agreement of standard uncertainty together with a second “e-mail” round with more steps. Stephan will try to get it ready by the end of March 2003 and in that case it can be included to the Solar Keymark final report.

###### Informative Annex about uncertainty

Kostas Voropoulos presented a suggestion of an informative Annex for uncertainty calculation. See

Annex WP1.A.B, “General guidelines for the assessment of uncertainty in solar collector efficiency testing” by Emmanouil Mathioulakis and Kostas Voropoulos.

Kostas requested for comments either directly or by e-mail correspondence. The discussion that followed concluded that the suggestion is an excellent base that explains the basic theory behind the calculations. Stephan Fischer explained that it is the same procedure as he use for the uncertainty Round Robin, where it actually will be tried out. However, it would be beneficial with a continuation with more direct recommendations on how to do, step by step, with figures. Stephan Fischer said that he could contribute with that when the uncertainty Round Robin is ready.  Pierre Richard said that he could contribute with input about the uncertainty at steady state conditions, which is a large source for the uncertainty. It was recommended that this work should continue and a ready suggestion should be prepared within the TC312 WG1 group.

Information about the German Round Robin

Harald Druck informed about the German Round Robin of solar collectors. Six laboratories are participating and each laboratory should measure the performance with the methods that they are accredited for. ITW are for example testing both with the QDT and the outdoor SS test. ITW is reference laboratory and will test the collectors both before and after the other laboratories. There are two collectors circulating in the Round Robin. The Round Robin will be finished in September 2003 and thereafter other countries are free to participate, which in principal will be the next season for other countries. In the end of March it will be decided how the results should be presented (confidential or not). It will be important to keep it confidential at least until other countries had time to do their testing. Harald will send a not to the Solar Keymark group when it is decided. For more information or interest in participating, please, contact Harald directly (drueck@itw.uni-stuttgart.de).

Preliminary investigation on procedures for an indoor exposure test

Amelie Veenstra informed that she had not been able to do any preliminary investigations due to lack of financing. It was recommended that this should be a work item for the TC312 WG1 group.

Proposal for change of test sequence

Pierre Richard informed about a suggestion for changing the stated test sequence in the standard, En 12975-2, 5.1. The main reasons for changing the test sequence are:

- The exposure test is not the same in all countries and it will therefore have different influence on the thermal performance test.

- The reliability tests are qualitative tests (expertise of labs) and therefore may give different results at different laboratories while the thermal performance test is quantitative. These different tests should be separated.

When changing the test sequence a few things should be considered:

- What happens with the collector during the reliability tests, specifically for innovative collectors there the test is not relevant.

- The test costs should be minimized.

- When do we have to use the sequence?

Pierre had several suggestions for how to change the test sequence:

- Allow using two collectors at the same time. One for thermal performance and one for reliability tests. This will decrease testing time.

- To do the thermal performance test both before and after the qualitative test. This will increase testing costs.

- To do the thermal performance test both before the qualitative test and then have an optional second thermal performance test again. This will not necessarily increase testing costs. For example could a manufacturer trust on the first result (for example if the insulation and the coating are well known) while the second test could be used for innovative products.

- To do thermal performance before the qualitative test.

- To change shall to should for the test sequence in the standard.

A discussion followed with the arguments that it is a reason for having the stated sequence. Primarily due to that the efficiency should be representative for a used product. It was decided that it should be recommended that this should be a work item for the TC312 WG1 group and that we need someone to test if the reliability test really is affecting the efficiency tests.

Benchmark test for parameter identification for QDT

Peter Kovacs presented the updated version of the inter-comparison of parameter identification from a set of measurement data. See Annex WP1.A.C, “Inter laboratory comparison on evaluation of collector test data” by Peter Kovacs.

The purpose with the benchmark is to compare if measurement values collected during a QDT test will be evaluated in the same way at different laboratories. In total seven laboratories took part in the comparison. Due to the large scatter in the first round a further and more detailed analysis was made with the following steps:

- Check of the regression algorithm itself with identical data. No differences were found.

- Calculations of solar incident angels.

- Calculation of specific heat and density of the heat transfer fluid.

This step revealed some discrepancy and as a result the following conclusion can be drawn:

- Several points in the standard text should be revised in order to:

    - make it easier to interpret,

    - make sure that the calculations will give the same result independent of laboratory.

- Thanks to this inter comparison have some laboratories changed their procedures. This will give a permanently improvement.

- This type of intercomparisons should precede a “real” Round Robin” with measurements in order to separate calculation errors from measurement errors.

The suggestions for improvements can be found in the document “Internal procedure for Solar Keymark WP1.a and input to CEN/TC 312”, Version 3, 2003-03-20, edited by Åsa Wahlström.

Not all laboratories participated and since it has been showed that the procedures in specific laboratories has been improved Peter welcomed the other laboratories to join the intercomparison. Preferably before May 2003 since he will conclude the intercomparison in a paper for the ISES 2003 conference.

Comments to the document WP1.A input to WG1 and discussions at WG1

There were no comments to the document that couldn’t be discussed in CEN/TC312 WG1 meeting. Since the WG1 meeting was informal no real decisions could be made and the discussions, conclusions and decisions are therefore further input from the Solar Keymark group to WG1 to be added in the third version of “Internal procedure for Solar Keymark WP1.a and input to CEN/TC 312”.

The discussed points are reported in Annex WP1.A.D, “Discussions, conclusions and decisions at the informal TC312/WG 1 meeting in Lisbon” edited by Åsa Wahlström.

ANNEX WP1A.A

Uncertainty power point presentation

by Stephan Fischer, ITW.

ANNEX WP1.A.B

General guidelines for the assessment of uncertainty in solar collector efficiency testing

by Emmanouil Mathioulakis and Kostas Voropoulos.

ANNEX WP1.A.C

Inter laboratory comparison on evaluation of collector test data

by Peter Kovacs.

|  | Solar Keymark |  |
| --- | --- | --- |

Peter Kovács

SP,  The Swedish National Testing and Research Institute

Box 857 501 15 Borås, Sweden

Tel: +46 33 16 56 62  Fax: +46 33 13 19 79  E-mail: peter.kovacs@sp.se

2003-02-26

### Summary

A dataset involving six days of measurements on a large flat plate collector was sent to the participants in the network of collector test institutes established during the Solar Keymark project. The measurements were made according to paragraph 6.3 of EN 12 975 (Quasi dynamic testing). A guide for the evaluation, describing the purpose of the exercise and the contents of the dataset and a general reporting format was also sent to the participants. In all, seven laboratories took part in the work and presented some end results. A previous issue of this report (2002-05-24) presented preliminary results from the investigation. In the present report, the investigation has been further elaborated in the respect that the calculation models were checked in more detail. The regression algorithm itself was checked using identical models and input data, but no differences were found. The calculations of solar incidence angle and of specific heat and density of the heat transfer fluid were also checked. These checks revealed some discrepancies and as a result, some of the participant’s end results have been updated, and the conclusions are further developed.

Through an iterative process, the accuracy of the final results (collector model parameters) have been substantially improved. As a step in the quality assurance process of the participating institutes, the calculation models and methodology have been permanently improved. However, to assure a permanent high level of accuracy similar to the one presented here, the standard’s text has to be revised so that it is easier to interpret. Suggestions of how it should be revised are presented in the conclusions.

### Aims and method

The first purpose of the comparison was to reveal any differences in the calculation models used at the different institutes. For this purpose, the participants made the required intermediate calculations and then performed the regression to the complete dataset in a first round. The collector model used is described in equation 1 below. These preliminary results (described in a previous issue of this report (2002-05-24)) revealed large discrepancies amongst the participants and therefore, the calculations were analysed in more detail in a second and third round. Here, the regression algorithm itself was checked using identical models and input data. The calculations of solar incidence angle and of specific heat and density of the heat transfer fluid were also checked.

(1)

Secondly, the aim was to reveal any unclear points concerning the data evaluation and selection in the standard. Several selection criteria’s are included in the standard. How were they interpreted? Did everyone find all the criteria?

In order to encourage a discussion about outliner treatment two datapoints were manipulated and the participants were instructed to treat outliners as they considered appropriate.

### Results

The results of the investigation can be summarized in figure 1a and 1b. An efficiency curve based on reference conditions according to 6.3.4.8.4 of EN 12975-2 shows good agreement for all seven parameter sets in figure 1b (within ( 0.3 % at tm-ta=50 (C). The improvement from the first try (figure 1a) is substantial.

| Fig. 1a. Efficiency curves based on parameters from seven reduced datasets. First results. | Fig. 1b. Efficiency curves based on parameters from seven reduced datasets. Final results. |
| --- | --- |

All parameters except c5 (figure 2 and equation 1), the thermal capacitance, showed similar good agreements when based on the reduced data sets (after exclusion criteria had been applied, appr. 600 data points in all). Main reasons for the large scatter in the first results were:

- Different interpretation of exclusion criteria (some differences still remain in final results)

- Errors in the models and incorrect reference temperatures used to calculate specific heat and density

- Errors in the models used to calculate incidence angle modifiers

Fig. 2 Parameter c5, the thermal capacitance was the parameter showing the largest variations when derived from the reduced datasets.

The reason for c5 being so low for participant 7 is that the term dtm/dt calculated from average tm-values was used in the regression and not the online calculated dtm/dt included in the test data file. The remaining scatter in c5 is probably due to different interpretations of the exclusion criteria related to temperature stability.

Looking at the parameters resulting from the unreduced  datasets (before exclusion criteria had been applied, 1642 data points in all) rather large differences among the participants were revealed. The heat loss coefficients (fig. 3) and the incidence angle modifiers for beam and diffuse radiation (fig 4a and b) from lab 4, 5 and 7 seem to be differing from the “main body” of results in a similar manner. One  explanation to this was that some labs made “corrections” to the dataset in order to have all inputs physically correct. So was e.g. beam irradiance set to 0 if, in a datapoint, diffuse irradiance was larger than global irradiance. Similarly the term (1/cos((I)-1) was set to zero for angles >85 degrees or set equal to the value at 80 degrees for angles larger than 80. Another reason was that some errors existed in the calculation models used for the calculation of incidence angles, or for the calculations of density and specific heat of the heat transfer fluid. The results based on the unreduced datasets were not updated during the course of the project i.e. the figures below show the results as they were after the first round, before errors and misinterpretations were corrected.

Fig. 3 Heat loss coefficients derived from unreduced datasets.

| Fig. 4a. F´()en and Kd derived from unreduced dataset | Fig. 4b. Kb derived from unreduced dataset |
| --- | --- |

Concerning the calculations of heat transfer fluid specific heat and density some significant differences were identified and the underlying errors or misunderstandings were corrected. After that, the scatter in both quantities stayed below ± 0.1 %. The differences in calculated incidence angles was, after correction of errors, also below ± 0.1 %.

The remaining number of data points after reduction according to the standards text is shown in figure 5. The spread has been reduced after some discussions around the first results. The remaining scatter is mainly due to:

- The Gd/G>0.5 criterion. According to the standard, these points should not be included in the data set. A should requirement is considered weaker than a shall requirement and thus, only one or two participants used this to exclude data.

- The dTinlet > ±1K criterion. Variations during a test sequence larger than this shall be excluded from the regression input data. Furthermore data recorded during “step-changes” should be excluded. Only two participants applied the former and one the latter criterion.

Fig. 5 Number of data points used in the regressions after selecting data according to the standard’s requirements.

The following selection criteria shall, according to the standard be used to exclude data.

- Tout- Tin <1(C

- G<300W/m2

- Tinlet not stable within ±1K during test day or test sequence

- Flowrate not stable within ±1% of the set value during test day or test sequence and may not vary more than ±10 % from the set value from one sequence to another.

The following selection criteria should, according to the standard be used to exclude data.

- Data recorded during step change

- Gdiff/G>0.5

- Power output<0 (which is a superfluous requirement since Tout- Tin <1(C and G<300W/m2)

A few selection criteria was “invented” from requirements on test data. If for example the inlet collector temperature in a sequence is to be chosen so that tm-ta lies within ± 3K around solar noon, it does not mean that the rest of the data from that sequence shall be excluded. Similarly, a requirement on average wind speed does not mean that data with wind speeds higher than 4 m/s or lower than 1 m/s shall be excluded.

The complete parameter sets based on unreduced and reduced data and diagrams showing the different parameter values are available in the excel-file “QDT testdata summary.xls”.

### Conclusions

- Some participants “corrects” measurement data on a routine basis in order to have all inputs physically correct and some don’t. This will in general have some effect on the final results and a check should therefore be indicated in the standard. Preferably in paragraph 6.3.4.6.4 “Evaluation of test data”. So should e.g. beam irradiance be set to 0 if, in a datapoint, diffuse irradiance is larger than global irradiance. Similarly the term (1/cos((I)-1) should be set equal to the value at 80 degrees for angles larger than 80.

- The selection criteria given in the standard are somewhat unclear and they are not efficiently organised in the document. Under paragraph 6.3.4.6.4 “Evaluation of test data” the shall-requirements for selecting data could be repeated/ compiled for clearness reasons. The should-requirement on power output >0 should be removed (se comment above). The should requirement on Gd/G>0.5 could probably also be removed (see note in paragraph 6.3.4.6.4).

- The online calculation of dtm/dt according to 6.3.4.5.2 shall be mandatory as it has a big impact on the final results if it’s calculated from average values instead.

- The properties of the heat transfer fluid shall be calculated as a function of collector average temperature (specific heat) and as a function of the temperature at the flowmeter (density) with a minimum accuracy of ± 0.1 % and not  ± 1 % as stated in the standard, paragraph 6.1.3.2.

- From the large majority of participants it was understood that outliners that cannot be explained shall not be excluded from the data set.

- There seem to be a few minor misunderstandings about sign conventions, units and number of significant figures in the parameter values reported.

- The standard deviations for each parameter resulting from the regression should be reported along with the parameter itself as this gives some additional information about the quality of the result.

ANNEX WP1.A.D

Discussions, conclusions and decisions at the informal TC312/WG 1 meeting in Lisbon.

Edited by Åsa Wahlström.

|  | Minutes of the informal TC312/WG1 meeting, INETI, Lisbon the 7th of March 2003 |  |
| --- | --- | --- |

Specific points in the document “CEN/TC312/WG1 –EN 12975-1&2 Amendment –List if incoming comments (03/02/2003)” (Ref. CEN/TC312/WG 1 N243) distributed by the WG1 Convener were discussed. Hereafter this document is called N243. The Convener has given each point a status, where:

C = comments foe which a consensus has been obtained in the Solar Keymark Project and in the CEN/TC312/WG1 Rome Meeting

NC = comments foe which no consensus has been obtained until now

N = New comment

Rain penetration test (En 12975 –1, 5.3.7 and EN 12975 –2, 5.7)

In the Solar Keymark group three new suggestions for the Rain Penetration test has been established as internal papers. Specific criteria from theses internal papers are mentioned in several points in N243. However it was decided that it should be recommended that the whole new suggestions should be included in the standard. In this way the carefully thought out wordings for each suggestion would also be used for improvement of the standard.

It has been concluded that the pass criteria for the Rain Penetration test, EN 12975 –1, 5.3.7, should be changed for the check by weighing the collector and by measurement of humidity N243 Point 5 and 6. In order to be consistent also the pass criteria for the measurement of the condensation level should be changed. The informal WG1 meeting decided that EN 12975 –1, 5.3.7, c, should be changed to “the measured condensation level shall be less than 10 % of the transparent cover”.

All points below are considering the N243 document points stated below comments to EN 12975-2 standard.

N243 Point 4;Inconsistence between 12975-1 and 12975-2

This point is noted NC in N243. However, this point has been examined and discussed by the Solar Keymark group. Excerpt from the minutes of the Solar Keymark Rome meeting:

Inconsistence between 12975-1 and 12975-2. Some requirements are subjective since description of test methods are missing. For example the requirement of  “no sharp edges” has no test method.

AAW informed that “no sharp edges” is the only requirement that she can find that has no test method.  She also informed that Hoang Liauw has checked with TC if it is a formal problem that the EN 12975-1 has requirements that have no method for testing in EN 12975-2. His answer is: “ There is no rule that a requirement in an European Standard should always refer to a test method in another accompanying European standard. It may well refer to an ISO or other reference document.”

Action: It was concluded that is not difficult to identify “ sharp edges” and the problem need no further actions.

### The informal WG1 meeting decided to recommend hat this point will not need any further discussions and should be deleted from N243.

N243 Point 13; Spraying of the collector in 5.7.2.1

### This point is noted C in N243. That “sprayed on all sides from above perpendicular to the horizontal” should be changed to “sprayed on all sides”.  However, it should not be changed to “sprayed on all sides” because this means that the collector also must be sprayed on the backside. This will only complicate the testing, with increased costs, without any improvements of the results. It is not raining on the back for a mounted collector. The informal WG1 meeting decided to recommend that “sprayed on all sides from above perpendicular to the horizontal” should be changed to “sprayed on exposed sides”.

# N243 Point 18; Special consideration for sensitive material in the collector backs, 5.7.2.2

### This point is noted C in N243. It recommends that special measures must be taken during the rain penetration spraying in order to not damage sensitive material. However, this point is only relevant if the spraying must be done on all sides. Since this should be changed according to the above previous point (N243, point 13) this point is not necessary anymore. The informal WG1 meeting decided to recommend hat this point will not need any further discussions and should be deleted from N243.

# N243 Point 22; Impact resistant test, 5.10.

This point is noted NC in N243. The impact test includes a non-realistic test. Although it is an optional test, this does give confusion on the market in practice. Also, in practice there was never reason to doubt the impact resistance of collectors. Solar Keymark has previous discussed if it would be better to delete this clause in the standard. The following discussions concluded that a test description, even if it is informal, is needed in the standard as several laboratories are actually using this test. It is important that those laboratories are using the same procedures. However, the test procedures needs improvements and a comparison on how these test are made for the window industry is needed. Hubert Fechner and Amelie Veenstra will investigate and give suggestions of improvements to the Convener. Continuing discussion concluded that it might be better to put the test as an informative Annex to clearly show that the test is optional as soon as the needed improvements has been made.

# N243 Point 24; Tilt angel fixed, 6.1.1.3.

This point is noted NC in N243 and has two suggestions for solutions. The Solar Keymark group has previously made one of the suggestions. Excerpt from notes from Solar Keymark Rome meeting:

##### Tilt angel of the collector

Problem: The tilt angel of the collector should be mounted in 45( according to clause 6.1.1.3. This will make comparison of measurements at different laboratories difficult since the incident angel will vary with the latitude.

Action: It was decided that it should be recommended in the “Solar Keymark input to WG1” that the same specifications as in Clause 6.1.4.3 should be added in the beginning of 6.1.1.3. The first two sentences on page 35: “ The angel of incidence of direct solar radiation at the collector aperture shall be in the range in which the incident angle modifier for the collector varies by no more than 2 % from its value at normal incidence. For single glazed flat plate collectors, this condition will usually be satisfied if the angle of incident of direct solar radiation at the collector aperture is less than 20(.”  Thereafter, should the word “shall” be changed to “should” in the first sentence in 6.1.1.3.  ( “ …..the collector should be mounted such …”)

### The other suggestion is that it is not necessarily to specify the tilt angle conditions since it will be specified by the conditions for the angle of incident of direct solar radiation declared in 6.1.4.3. The informal WG1 meeting decided to recommend hat the first proposal should be used together with a change for the first sentence in 6.1.1.3 to “the tilt angel of the collector should be mounted in 45(+/- 10 (”. The mounted tilt angel should also be stated in the test report.

# N243 Point 50; Specification of surrounding air speed, 6.3.4.3.

This point is noted NC in N243. The question is if the surrounding air speed should be greater than 1 m/s and less than 4 m/s or only less than 4 m/s. The discussion continued with that it might be better to have the speed specified to 4 +/- 1 m/s or 3 +/- 1 m/s as in the SS test. Peter Kovacs and Stephan Fischer will investigate this further and report the result with recommendations to the Convener.

N243 Point 54; Remove the requirement of Gd/G>0.5, 6.3.4.3.

### This point is noted NC in N243. The discussions concluded that it could be removed and still be competitive with SS. The informal WG1 meeting decided to recommend hat this point will be considered as C and do not need any further discussions.

N243 Point 60; The test report have to be brought in accordance with the requirements of EN ISO 17025, Annex D.

This point is noted C in N243, but the recommendation for changes is only “Modify Annex D”. It is a need for a further explanation on how the Annex should be modified in order to be in accordance with EN ISO 17025. Stephan Fischer will investigate this further and report the result with recommendations to the Convener.

N243 Point 62; Addition of an informative Annex for calculating the uncertainty.

This point is noted C in N243. The discussions about this concluded that the Document provided by “Demokritos” needs further work as stated during the Solar Keymark meeting on the 5th of March. Further input are requested:

- Input by email to Demokritos from other laboratories

- More direct recommendations on how to do, step by step, with figures. Based on the results from the uncertainty Round Robin made by Stephan Fischer

- Input by Pierre Richard about the uncertainty at steady state conditions.

Another discussion was if this document really is needed in the standard since it only explains things that already are in the standards and in other standards. However the explanations are made in a much more comprehensive and understandable way. Maybe it should be a Solar Keymark document instead.

N243 Point 11; Need of an inodor exposure test, 5.4.

This point is noted C in N243. Jan Erik Nielsen will apply for financing to developing such a test within the EC sixth framework program.

N243 Point 56; Other method than MLR, 6.3.4.8.1.

This point is noted C input from ITW in N243. Stephan Fischer will investigate this further and report the result with recommendations to the Convener.

N243 Point 59; For how long and for how many datapoints should the value exceed

+/- 0.005 K/s, 6.3.5.2.

This point is noted further discussions in N243. Stephan Fischer will investigate this further and report the result with recommendation to the Convener.

N243 Point 55; Other method than MLR, 6.3.4.8.1.

This point is noted C input from ITW in N243. Stephan Fischer will investigate this further and report the result with recommendation to the Convener.

Other inputs: Annex L needs an equation instead of a table.

Stephan Fischer will provide the equation to the Convener.

Other inputs: Change of test sequence, 5.1.

The discussions about this made on the Solar Keymark meeting on the 5th of March will be summarised by Åsa Wahlström and sent in the third version of “Internal procedure for Solar Keymark WP1.a and input to CEN/TC 312” to the Convener.

# Text boxes

Fel! Ogiltig länk.

Fel! Ogiltig länk.